## Slide 1: Photoacoustic and Ultrasonic Tomography for Breast Imaging
Felix Lucka*, Ben Cox &
the PAMMOTH team (*he/him, felix.lucka@cwi.nl)
Applied Inverse Problems, 7th Sep 2023
Notes: TODO: New high res image

## Slide 2: Breast Cancer
Most common cause of cancer death in women worldwide:
25% of all cancer cases
15% of all cancer deaths
Early detection and diagnosis are key to successful treatment!

## Slide 3: Conventional Breast Imaging
Magnetic resonance imaging
contrast agent
expensive
availability
not suitable for every woman
Ultrasonography
whole breast imaging is difficult
operator dependent
courtesy of Rianne Bulthuis and Maura Dantuma
X-ray mammography
painful compression
ionizing radiation (tumor induction)
glandular tissue is problematic
Urgent need for novel imaging techniques providing higher specificity, contrast and image resolution than X-ray mammography at lower costs than MRI.

## Slide 4: Photoacoustic Tomography (PAT)
adapted from Rianne Bulthuis and Maura Dantuma
Hybrid imaging: light in, sound out
non-ionizing, near-infrared radiation
quantitative images of optical properties
novel diagnostic information

## Slide 5: Photoacoustic Imaging: Spectral Properties
Oxyhemoglobin
Deoxyhemoglobin
Elastin
Collagen
Lipid (a) (b)
Water
(from Beard, 2011)
different wavelengths allow quantitative spectroscopic examinations
gap between oxygenated and deoxygenated blood
use of contrast agents for molecular imaging

## Slide 6: Ultrasound Tomography (UST)
adapted from Rianne Bulthuis and Maura Dantuma
sound in, sound out
different from conventional US but as safe
quantitative images of acoustic properties
novel diagnostic information

## Slide 7: Speed of Sound vs MRI Images
From: Duric, Littrup, 2017. Breast Ultrasound Tomography, IntechOpen

## Slide 8: H2020: PAMMOTH
Combined PAT+UST scanner to obtain novel diagnostic information from high resolution maps of optical and acoustic properties

## Slide 9: PAM3 System

## Slide 10: Our Contributions
simulation studies for
ultrasonic transducer specification
light excitation design
sensing pattern design
measurement protocol design
reconstruction algorithm design
accuracy vs. computational time/resources/complexity
scanner modelling
assist high performance computing implementation
assist phantom & calibration design
process data, refine measurement procedures

## Slide 11: PAT: Mathematical Modelling
radiative transfer equation (RTE)
photoacoustic effect
acoustic wave equation
measurement (on boundary)
two coupled inverse problems:
acoustic initial value problem with boundary data.
optical parameter identification with internal data.

## Slide 12: Reconstruction of Initial Photoacoustic Pressure
linear inverse problem
variational approach
first order optimization with early stopping
model acoustic properties, model discrepancies
model /calibrate piezoelectric sensor properties: sensitivity, impulse response, angular sensitivity, …
parameter choices, image artifacts, …
high res (0.4mm), broadband numerical wave simulations

## Slide 13: PAT Modeling, baseline
Notes: SOS model: cup

## Slide 14: PAT Modeling, + spatial response
Notes: SOS model: cup

## Slide 15: PAT Modeling, + temporal response
Notes: SOS model: cup

## Slide 16: PAT Modeling, + sampling density
Notes: SOS model: cup

## Slide 17: PAT Modeling, higher resolution
Notes: SOS model: cup

## Slide 18: Illustration: Phantom Validation

## Slide 19: Illustration: Phantom Validation

## Slide 20: Illustration: In Vivo Results
>1
<1
Intensity ratio 890 nm / 720 nm
Figure from: Spectrum Medical, Non-Invasive Measurement of SaO2, SvO2 & Hemoglobin

## Slide 21: How deep can we image?
Increasing depth from skin
medial-lateral depth color-coded PAT MIPs
48 mm

## Slide 22: UST: Mathematical Modelling (simplified)
acoustic wave equation
measurement (on boundary)
multiple sources
inverse problem:
acoustic parameter identification problem with boundary data.

## Slide 23: UST Reconstruction Approaches
Travel time tomography: geometrical optics approximation
robust & computationally efficient (time-of-flight picking followed by linear inversion)
valid for high frequencies (attenuation!), low res, lots of data
Javaherian, L, Cox, 2020. Refraction-corrected ray-based inversion for three-dimensional ultrasound tomography of the breast. Inverse Problems
Full waveform inversion (FWI): fit full wave model to all data
high res from little data, transducer modelling, constraints
many wave simulations, complex numerical optimization
lower TRL but already used in 2D systems
our system is not designed for it

## Slide 24: Time Domain Full Waveform Inversion
Pérez-Liva, Herraiz, Udías, Miller, Cox, Treeby, 2017. Time domain reconstruction of sound speed and attenuation in ultrasound computed tomography using full wave inversion. JASA
gradient for first-order optimization via adjoint state method
where is the time-reversed data discrepancy
two wave simulations for one gradient (for one source )
Starting point in 2D:
, s.t.

## Slide 25: 3D Time Domain FWI for Breast UST
Challenges and solutions for 3D:
2 x wave simulations per gradient  stochastic quasi-newton optimization (SL-BFGS)
computationally & stochastically efficient gradient estimator  source encoding for time-invariant systems
memory requirements of gradient computation  time-reversal based gradient computation
slow convergence and local minima  coarse-to-fine multigrid schemes
computational resources  runs on single GPU, can utilize multiple GPUs
,

## Slide 26: Numerical proof-of-concept study
L, Pérez-Liva, Treeby, Cox, 2021. High Resolution 3D Ultrasonic Breast Imaging by Time-Domain Full Waveform Inversion. Inverse Problems

## Slide 27: Numerical proof-of-concept study
L, Pérez-Liva, Treeby, Cox, 2021. High Resolution 3D Ultrasonic Breast Imaging by Time-Domain Full Waveform Inversion. Inverse Problems
Starting point
Improvement

## Slide 28: FWI for Experimental Data: Where Are We?
data from phantom objects, volunteers & patients
ray-based travel time tomography reconstructions
photoacoustic reconstructions  data pre-processing, scanner & transducer modelling, wave simulations
modeling of US protocol, data read-in & pre-processing
model calibration
FWI of phantom objects, quantitative evaluation
FWI of volunteer / patient data
Clinical evaluation

## Slide 29: SOS Phantom
X
Z
SOS – cwater
-100 0 +100
Y
Notes: 50.000 source activations

## Slide 30: SOS Phantom: Travel Time Tomography
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
2.0mm grid; (bent-ray)
SOS – cwater

## Slide 31: SOS Phantom: FWI
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
2.0 mm grid, 0.37/0.21 MHz, 0.22% original signal
SOS – cwater

## Slide 32: SOS Phantom: FWI
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
SOS – cwater
1.5 mm grid, 0.50/0.28 MHz, 0.52% original signal

## Slide 33: SOS Phantom: FWI
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
SOS – cwater
1.0 mm grid, 0.75/0.42 MHz, 2.47% original signal

## Slide 34: SOS Phantom: FWI
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
SOS – cwater
0.8 mm grid, 0.94/0.53 MHz, 6.55% original signal

## Slide 35: SOS Phantom: FWI
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
SOS – cwater
0.6 mm grid, 1.25/0.71 MHz, 21.35% original signal

## Slide 36: In Vivo Results: TTT
-100 0 +100
X
Z
SOS – cwater
Y
2.0mm grid
bent ray

## Slide 37: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
2.0mm grid
0.22% original signal
Y

## Slide 38: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
Y
1.5mm grid
0.52% original signal

## Slide 39: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
Y
1.0mm grid
2.47% original signal

## Slide 40: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
Y
0.8mm grid
6.55% original signal

## Slide 41: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
Y
0.6mm grid
21.35% original signal

## Slide 42: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
Y
0.4mm grid
65.46% original signal
PAT shows that these three dots are blood vessels

## Slide 43: Improvement (up to now…)
Y
X
Z
SOS – cwater
Y
X
Z
SOS – cwater
Y
-100 0 +100

## Slide 44
Y
-100 0 +100
X
Z
SOS – cwater
Y
X
Z
SOS – cwater
Y
Improvement (up to now…)

## Slide 45: Using 2SOS model in PAT
Notes: Colour codes depth, everything above 3cm is clipped off

## Slide 46: Using TOF model in PAT
Notes: Colour codes depth, everything above 3cm is clipped off

## Slide 47: Using FWI model in PAT
Notes: Colour codes depth, everything above 3cm is clipped off

## Slide 48: Summary & Outlook
need for novel breast imaging techniques
photoacoustic (PAT) and ultrasound tomography (UST) give access to high-quality images of optical and acoustic tissue properties
combined PAT+UST scanner designed & build
evaluation on data from phantoms, volunteers & patients
PAT reconstructions work well
proof-of-concept studies of FWI for high resolution 3D UST
demonstration of FWI on experimental data
optical inversion in PAT (QPAT)
L, Pérez-Liva, Treeby, Cox, 2021. High Resolution 3D Ultrasonic Breast Imaging by Time-Domain Full Waveform Inversion. Inverse Problems
Felix.Lucka@cwi.nl Photoacoustic and Ultrasonic Tomography for Breast Imaging
Dantuma et al., 2023. Fully three-dimensional sound speed-corrected multi-wavelength photoacoustic breast tomography. arXiv:2308.06754